## Slide 1
Core Facility for Cell Sorting and Cell Analysis
What is wrong with my flow cytometry data?
Hints, tricks and pitfalls

## Slide 2
2
Webinar 4
02.05.2023
Compensation and unmixing:
Discover errors and correct them

## Slide 3
Core Facility for Cell sorting and Cell analysis
Everything has to do with spectral overlap…
Fluorescence dyes are NOT excited by only one laser and do NOT have a sharp emission peak
Emission peaks do overlap
https://www.miltenyibiotec.com/DE-en/resources/macs-academy/on-demand-webinars-and-tutorials.html#

## Slide 4
Core Facility for Cell sorting and Cell analysis
How to correct for spectral overlap…
Compensation
  process to correct for spectral spillover of one fluorescence dye to neighbouring detection channels
Unmixing
  calculates the contribution of each known fluorophore’s spectra to the total collected emission signal
Empirically calculated from single stain controls in multi-color experiments
‚false‘ signals

## Slide 5
https://www.miltenyibiotec.com/DE-en/resources/macs-academy/on-demand-webinars-and-tutorials.html#
Compensation…more or less?

## Slide 6
https://www.miltenyibiotec.com/DE-en/resources/macs-academy/on-demand-webinars-and-tutorials.html#
spillover
spread
Compensation…more or less?

## Slide 7: Spillover spread…is?
spread: mathematical introduced because of compensation/unmixing
brighter dyes  bigger spread
depends on dyes  important for panel design
spillover
spread
https://www.researchgate.net/figure/Spreading-error-and-loss-of-detection-sensitivity-a-APC-data-spread-into-the-Alexa_fig2_310494404

## Slide 8
Guidelines for the use of flow cytometry and cell sorting in immunological studies (third edition) - Scientific Figure on ResearchGate. Available from: https://www.researchgate.net/figure/Spreading-error-and-fluorochrome-brightness-in-panel-design-and-common-compensation_fig4_356873359 [accessed 30 Apr, 2023]
Spillover spread…is?

## Slide 9: Golden rules for compensation/reference controls
Ctrls must be stained with only ONE fluorophor
Ctrls need to be as bright or even brighter as the samples fully stained
Ctrls must contain a negative and a positive population (unstained)
Background (autofluorescence) of positive and negative population needs to be the same
 compare cells with cells, beads with beads
Ctrl dye must fit experimental color  FITC ≠ GFP ≠ AF488
Beads good for rare cell markers or use higher expressed markers
Tandem-dyes differ from lot to lot
Buffers, fixation, … may change the fluorophors  changes in compensation

## Slide 10: Golden rules for compensation/reference controls

## Slide 11: Compensation Matrix in FlowJo
FlowJo allows different ways for compensation
  to start from scratch
  to use assisted calculation
  edit matrix from aquisition
Special case – the BD FACS Calibur
  if compensated during acquisition data will be saved
   no way back to the uncompensated data
  BUT post-acq editing via FlowJo

## Slide 12: Compensation Matrix in FlowJo

## Slide 13: Post-acquisition unmixing in SpectroFlo

## Slide 14: Compensation/Unmixing: Take home message
Compensation/unmixing of multi-color experiment is calculated from single stain controls
Single stained ctrl as bright or even brighter than sample
Use alternative markers/beads when signal is dim
Titrate antibodies
Autofluorescence of positive and negative fraction for each marker needs to be the same
Spread is introduced and can’t be avoided (pay attention to expert panel design)
FlowJo allows to view/edit or create new compensation matrices

## Slide 15: Thank you for your attention. See you next month: 6th June Next topic The alpha and omega of preparing cells for flow cytometry